## Slide 1: Geometric (Lens) Calibration
Universal lens calibration 장비 | 알고리즘 순서도
Lens Calibration 결과 전/후 | 
overview
<왜곡 보정 전>
<왜곡 보정 후>
LED dots.
25 mm 간격.
940 nm
LiDAR 안착 지점
3축 회전 모터
3축 Rail
LED dots.
50 mm 간격.
940 nm
Are the blobs detected?
Image is distorted.
Detect the center of the blobs.
Yes
No
Calculate camera parameters
End
Are there more than 4 blobs in the LED chart?
Image upscaling using FFT.
Redefine the basis vector using PCA.
Take more than 3 images of the chart.
Sort the points
Are the blobs sorted?
Yes
No
Yes
No
Prepare another LED chart
Are the images undistorted?
Undistort the images
Yes
No
Zhang’s technique
New process for the Sequential Flash LiDAR
Downscale the obtained parameters
Obtain new camera matrix to recover the FOV.

## Slide 2: Geometric (Lens) Calibration
Blob detection 결과 | Principle component analysis (PCA) 분석 결과
 | 곡률 계산
상세 설명
<The upscaled image>
100개 탐지
<The raw image>
3개 탐지
④ LED dot sorting
① Random dot detection
② PCA를 이용해 basis vector 계산
② basis vector rotation
⑤ basis vector 원복
보정 전 곡률:
보정 후 곡률:
정면 기준, 곡률 94% 감소
곡률 정의 는 접선의 기울기.

## Slide 3
왜곡 보정 후 FOV 감소 확인 | 원인 분석
감소량 측정 결과 | 
Geometric (Lens) Calibration
향후 계획
675 mm
417 mm
425 mm
시야각 약
시야각 약
675 mm
(여분을 고려했을 때 )
417 mm
425 mm
Raw image
왜곡 보정
왜곡 보정 전
왜곡 보정 후
왜곡 보정 후 외각 부분이 늘어나며 FOV 바깥으로 밀려나게 됨.
FOV 를 회복할 수 있는 방안 모색 중.
왜곡 보정 후 얻어야 할 영역
왜곡 보정 후 현재 얻고 있는 영역
(crop 된 영역)