## Slide 1
Falcon/Eagle LGIT | Tesseract
System Overview
Camera calibration tool (Single mode green laser source +
2D beam splitting diffractive optical instrument)
532nm (green) nominal wavelength
Laser safety Class 1 compliant
Illuminator optics project an infinite conjugate spot pattern
Standard calibration outputs:
Intrinsic parameters:
  • Practical Focal Length (PFL) in x and y, units of pixels
  • Principal point (PP) in x and y, units of pixels
  • Distortion polynomial coefficients
Extrinsic parameters
• 3 Degree of Freedom (DOF) angular pose of camera relative to Illuminator.
Required initial conditions for model solver
  • Nominal focal length (pixels)
Hardware Quartus will provide:
4-Channel Laser Module
Illuminator module
Single mode fiber patch cable
Hardware LGIT will provide:
Socket
Frame grabber
Fixture to move socket to illuminator
Enclosure (optional)
Illuminator
Optic fiber
Laser Module
DUT
Socket
F/G
Moving fixture
Illuminator
Mounting Plate
DUT
WD
FOV coverage > 130°
DUT FOV
Illuminator description:
50mm clear aperture
Diffractive pattern will cover > 130° DFOV
WD depends on module FOV
Model | Eagle (NFOV) | Falcon (MFOV)
Expected Max WD (mm) | >50mm | ~23mm
Notes: 과거 이력, 실패 이유 등 확인

## Slide 2
Falcon/Eagle LGIT | Tesseract
Hardware Details
Laser module (PN: Q5179-0160)
Net weight: 8.9kg
Dimension: 438mm X 380mm X 145mm
Specs:
  - Supply 4 channels of light output
  - 532nm (green light)
  - FC/APC fiber output
  - AC voltage input (100-250 VAC, 50/60Hz), 50W max power
  - USB port for SW control and calibration storage
  - Interlock input for safety switch
  - Indicator LED
Accessories included: power cord, USB cord
Tesseract XLHD Illuminator (PN: Q5179-0050)
Net weight: 3.6kg
Dimension: 420mm X 185mm X 134mm
Specs:
  - Optical test pattern generator for camera calibration
  - FC/APC fiber output
  - USB port for calibration storage
  - Remove shipping cap after installed in fixture
Accessories included: FC/APC fiber optic patch cable, USB-C cord
Notes: 과거 이력, 실패 이유 등 확인

## Slide 4
Illuminator
DUT
Illuminator
DUT
Horizontal
Vertical